## Lot 1 (hidden)
Please seperate different tabs for each compliance matrix
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 1 - Simulator
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | International Electrotechnical Commission, Medical Electrical Equipment, Part 1: General requirements for Safety, Rep. IEC 601-1, IEC, Geneva (1988)
2.2 | International Electrotechnical Commission, Medical Electrical Equipment, Part 2-29: Particular requirements for the safety of radiotherapy simulators, Rep. IEC 601-2-29, IEC, Geneva (1999)
2.3 | International Electrotechnical Commission, Radiotherapy simulators: Guidelines for functional performance characteristics, Rep. IEC 61170, IEC, Geneva (1993)
2.4 | “Commissioning of Radiotherapy Treatment planning systems: Testing for typical external beam treatment techniques”, IAEA, Vienna (2008) (http://www-pub.iaea.org/MTCD/Publications/PDF/te_1583_web.pdf)
2.5 | “Setting up a Radiotherapy Programme: Clinical, Medical Physics, Radiation Protection and Safety Aspects”, IAEA, Vienna (2008) (http://www-pub.iaea.org/MTCD/publications/PDF/pub1296_web.pdf)
2.6 | “Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards”, IAEA, Vienna (2014) (http://www-pub.iaea.org/MTCD/Publications/PDF/Pub1578_web-57265295.pdf)
2.7 | All the applicable International Atomic Energy Agency Safety Standards.
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | The System shall include an isocentric digital conventional simulator with local and a remote workstation for offline processing, design and approval of field parameters and shielding. Secure connectivity to the End-User intranet to facilitate data and image transfer is a requirement, as well as the ability to communicate with the Site’s PACS system for archiving.
4.1.2 | The fluoroscopy based unit must be capable of radiotherapy simulation to currently acceptable standards of radiotherapy practice.
4.2 | Technical Requirements
4.2.1 | The gantry shall have the following characteristics: Motorized gantry with isocentric design capable of 0 - 3600 rotation, an adjustable X ray focus to isocentre distance of 80–120 cm, an isocentre height above floor level ≤ 130 cm and an isocentre maximum sphere diameter of 3.0 mm
4.2.2 | The couch table must be entirely X ray transparent and graduated in all typical treatment couch movements, which can be related directly to those of the treatment unit.
4.2.3 | Independent collimation capabilities.
4.2.4 | A 50 kW high frequency generator (40 – 130 kVp).
4.2.5 | Backup Xray collimators shall provide auto-track to the size of the imaging device.
4.2.6 | Imaging system consisting of at least a 12” image intensifier, including in room and console image displays.
4.2.7 | An audible alert shall sound for fluoroscopy times in excess of 5 minutes.
4.2.8 | A digital flat panel detector of minimum size 40 cm x 40 cm.
4.2.9 | Two lead aprons for in-room presence during fluoroscopy procedures.
4.2.10 | A CCTV and a patient intercommunication system
4.2.11 | An in room monitor displaying the setup parameters.
4.2.12 | Control panel capable of remote mechanical setup of the device. Hand and couch control for local setup.
4.2.13 | Anti-collision warning devices, access interlocks, safety and radiation warning lights
4.2.14 | Wall-mounted laser system indicating the position of the isocentre including 2 side lasers (vertical and longitudinal) and 1 sagittal laser.
4.2.15 | Local (control console) and remote (treatment planning area) simulation review workstations for image processing, comparison, field design and approval
4.2.16 | Direct connectivity of the simulation workstations to the brachytherapy and external beam treatment planning systems (models tbc), virtual simulation workstation (model tbc) and record and verify system (model tbc) in DICOM (RT) format to facilitate exchange and transfer of images, machine parameters (including multi-leaf collimator settings) as well as digitally reconstructed radiographs.
4.2.17 | Direct connectivity to the Site’s PACs system for long-term image archive.
4.2.18 | External hard drive backup system and 30 minutes UPS to allow controlled shut down in the event of a power failure.
4.2.19 | The System shall form part of the End-User’s intranet and shall not have access to the internet. Software installation and upgrades/updates shall be performed using cd/dvd and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection.
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
15 | Uptime and Penalties
15.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
15.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
15.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
15.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 2 (hidden)
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 2 - CT Simulator
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | International Electrotechnical Commission, Medical Electrical Equipment, Part 1: General requirements for Safety, Rep. IEC 601-1, IEC, Geneva (1988)
2.2 | International Electrotechnical Commission, Medical Electrical Equipment, Part 2-29: Particular requirements for the safety of radiotherapy simulators, Rep. IEC 601-2-29, IEC, Geneva (1999)
2.3 | International Electrotechnical Commission, Radiotherapy simulators: Guidelines for functional performance characteristics, Rep. IEC 61170, IEC, Geneva (1993)
2.4 | “Commissioning of Radiotherapy Treatment planning systems: Testing for typical external beam treatment techniques”, IAEA, Vienna (2008) (http://www-pub.iaea.org/MTCD/Publications/PDF/te_1583_web.pdf)
2.5 | “Setting up a Radiotherapy Programme: Clinical, Medical Physics, Radiation Protection and Safety Aspects”, IAEA, Vienna (2008) (http://www-pub.iaea.org/MTCD/publications/PDF/pub1296_web.pdf)
2.6 | “Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards”, IAEA, Vienna (2014) (http://www-pub.iaea.org/MTCD/Publications/PDF/Pub1578_web-57265295.pdf)
2.7 | All the applicable International Atomic Energy Agency Safety Standards.
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | The simulator shall be a CT scanner for radiotherapy treatment simulation.
4.1.2 | The CT scanner shall have a carbon fibre flat table-top with indexing facilities (for all kinds of immobilization systems used in radiotherapy) identical to that of Linac (model tbc) and conventional simulator (model tbc) at the Site.
4.1.3 | The CT scanner shall have conventional in-built lasers or light beams, which indicate the coincidence of the centre of rotation and scan position.
4.1.4 | A second set of external lasers (ceiling and wall mounted) shall be installed in order to mark and setup the patient on the table-top distant from the gantry structure.
4.1.5 | The System must be networked to the external beam and brachytherapy treatment planning systems (models tbc), the record and verify system (model tbc), the simulation review workstation (model tbc), the laser imager and the End-User’s Site PAC system (for long-term archiving of images). DICOM (RT) format shall be used.
4.2 | Technical Requirements
4.2.1 | CT Scanner: The CT scanner shall be a whole-body spiral CT scanner;
4.2.2 | Gantry:
4.2.2.1 | A gantry aperture of at least 80 cm.
4.2.2.2 | A gantry tilt of at least 150
4.2.2.3 | Scan field of view of at least 50 cm.
4.2.2.4 | The couch positioning indicators in the gantry must have a positioning accuracy of ± 1 mm or better;
4.2.3 | Couch:
4.2.3.1 | The couch top material must be carbon fibre, flat bed type, with minimum dimensions of 235 cm x 40 cm, having a longitudinal moving range of 170 cm or more.
4.2.3.2 | The scannable range shall be at least 150 cm. It must be able to take a maximum weight of 200 kg or more without any change in stated performance specifications.
4.2.4 | X-ray System:
4.2.4.1 | High frequency X-ray generator with power rating of at least 50 kW (broad focus). kVp shall be in the range of 80 kV to 140 kV.
4.2.4.2 | The mA range from 30 mA to 400 mA, with step size of 5 mA or better. Peak anode heat dissipation rate of at least 800 kHU / min or better.
4.2.4.3 | The X-ray tube shall have dual focal spots.
4.2.5 | Detectors: The detectors shall be a high performance, low noise, high data density, active response data acquisition system. The detectors shall be solid state and free from repeated calibrations;
4.2.6 | Scan parameters:
4.2.6.1 | Slice thickness/spacing shall be at least 1 – 10 mm. Retrospective reconstruction shall be possible on raw data files with change in parameters such as FOV.
4.2.6.2 | The following scanning modes shall be possible: SPR (scout views), axial and spiral. The SPR length shall be more than 1500 mm long and the width must be at least 500 mm. It must be possible to obtain the SPR from AP or PA or lateral directions. The ability of the system to reproduce the scanning protocol from the SPR shall be ± 3 mm or better. The accuracy of distance measurements in the SPR (taken at isocenter distance) must be better than twice the pixel dimension.
4.2.6.3 | Different selection of pitch shall be possible, from 0.5 to 3. State the pitch available.
4.2.6.4 | The System must have automatic mA control software that automatically adjusts mA for patient size, adjusts mA along the z-axis, and modulates mA during rotation.
4.2.6.5 | Typical scan protocols for routine scanning (adult and paediatric) and quality control shall be available.
4.2.7 | Image Quality:
4.2.7.1 | The minimum reconstruction matrix must be at least 256 x 256 or higher.
4.2.7.2 | High contrast spatial resolution: at least 15 lp/cm maximum at 0% modulation transfer function (MTF). Low contrast detectability: 5 mm or less @ 0.3% using a Catphan 600 phantom (will be provided) on 10 mm slice thickness. The CT number accuracy must be better than 0 ± 4 HU for water and -1000 ± 5 HU for air;
4.2.8 | Computer System for the CT scanner: A high end computer system must be provided.
4.2.8.1 | There must be two monitors at the console and they must be 19" or more TFT flat screen LCD monitors. One of these will be used for acquisition and the other will be used for image review/basic processing.
4.2.8.2 | The hard disk capacity of the main computer system shall be at least 100 GB. The maximum possible hard disk capacity must be provided.
4.2.8.3 | For local copies of individual studies or software upgrade, a DVD reader/writer shall be provided
4.2.8.4 | The System shall be fully DICOM compliant. The System shall support the following: DICOM 3.0 Print service class as a user, DICOM 3.0 Storage class as a user, DICOM 3.0 Storage class as a provider, DICOM 3.0 Send / Receive, DICOM 3.0 Query / Retrieve service class as a user and DICOM 3.0 Query / Retrieve service class as a provider. DICOM compliance statement shall be provided;
4.2.8.5 | CT images shall be directly transferrable to the virtual simulation workstation, external beam and brachytherapy treatment planning systems (models tbc), and the Site PACS.
4.2.8.6 | A networked dry laser imager shall be provided (multiformat)
4.2.8.7 | A virtual simulation workstation shall be provided for volume definition, isocentre localisation, field placement and field design (shielding blocks or MLC) of external beam radiotherapy fields (see Appendix 1 for installation configuration), including generation of digitally reconstructed radiographs. It shall be possible to seamlessly export patient administrative data, images, volumes and machine parameters to the linac (model tbc), external beam treatment planning system (model tbc), simulator workstation (model tbc), laser imager and the record and verify system (model tbc).
4.2.8.8 | The workstation shall be equipped with an automatic overnight backup system to an external hard drive.
4.2.8.9 | Direct connectivity to the Site PACs system is required for long-term image archive.
4.2.8.10 | The system shall form part of the End-User’s intranet and shall not have access to the internet. Software installation and upgrades/updates shall be performed using CD/DVD and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection.
4.2.9 | Safety:
A bi-directional speaker communication must be provided between the operator and the patient;
The console shall be positioned such that the operator has full view of the patient through the lead glass window
A CCTV system shall be available to view the patient from the back of the gantry
Radiation warning lights and a door interlock shall be installed.
4.3 | Essential accessories shall be included with the System
4.3.1 | UPS: On line UPSs to for the console and virtual simulation workstation for at least thirty minutes in order to allow shutdown in the event of a power failure;
4.3.2 | Dose computation & display: The system shall display CTDI and DLP;
4.3.3 | Two lead aprons
4.3.4 | Injector for contrast administration operated from the console.
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
15 | Uptime and Penalties
15.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
15.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
15.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
15.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 3 (hidden)
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 3 - Medical Linear Accelerator with Record and Verify System
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | INTERNATIONAL ATOMIC ENERGY AGENCY "Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards, General Safety Requirements Part 3", IAEA, Vienna (2014)
2.2 | INTERNATIONAL ATOMIC ENERGY AGENCY, “Setting up a Radiotherapy Programme: clinical, medical physics, radiation protection & safety aspects”, IAEA, Vienna (2008)
2.3 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Medical Electrical Equipment, Part 1: General requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012, IEC, Geneva (2012)
2.4 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Medical Electrical Equipment, Part 60601-2-1: Particular Requirements for the Basic Safety and Essential Performance of Electron Accelerators in the Range 1 MeV to 50 MeV, Rep. IEC 60601-2-1:2009+AMD1:2014, IEC, Geneva (2014)
2.5 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Medical electrical equipment - Medical electron accelerators - Functional performance characteristics, IEC 60976, IEC, Geneva (2007)
2.6 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Radiotherapy equipment - Coordinates, movements and scales, IEC 61217, IEC, Geneva (2011).
2.7 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Medical Electrical Equipment: Requirements for the safety of radiotherapy treatment planning systems, Rep. IEC 62083:2009, IEC, Geneva (2009)
2.8 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, General requirements for safety 4. Collateral standard: Programmable electrical systems, Rep. IEC 60601-1-4 Part 1, IEC, Geneva (1996)
2.9 | International Electrotechnical Commission (IEC) developed IEC 62274 ed.1.0, ‘Safety of Radiotherapy Record and Verify Systems’ [20]. To quote from the IEC 62274 ed.1.0
2.10 | IAEA Human Health Report No. 7, ‘Record and Verify Systems for Radiation Treatment of Cancer: Acceptance Testing, Commissioning and Quality Control’ (2013).
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | Be capable of producing three photon beam energies (namely 4 MV, 6MV and 10/15 MV) and four electron beam energies (in the range 5 – 15 MeV) for treating cancer patients with external beam radiotherapy;
4.1.2 | Be equipped with an EPID and an MLC;
4.1.3 | Be capable of treating 40 – 60 patients per day per shift using 2D and 3D CRT techniques;
4.1.4 | Be capable of (future) upgrade to IMRT technology;
4.1.5 | Be supplied with a RVS including a server and workstations, that forms an intranet that links the imaging, dosimetry and treatment planning equipment at the End-User;
4.2 | Technical Requirements
4.2.1 | Linac:
4.2.1.1 | Photon energies 4, 6 and 10/15 MV all with flattening filter;
4.2.1.2 | Four electron energies from 5 – 15 MeV
4.2.1.3 | IEC 61217 scale convention in clinical mode;
4.2.1.4 | A motorized gantry with isocentric design, 100 cm source to axis distance and gantry rotation ± 185°;
4.2.1.5 | The mechanical isocentre shall have a maximum diameter of ≤ 2 mm for all three rotation axes (collimator, gantry, EPID imaging system and treatment table);
4.2.1.6 | Collimator rotation at least ± 100° with motorized rotation;
4.2.1.7 | Asymmetric jaw movements;
4.2.1.8 | Optical distance indicator with range at least SAD = ±30 cm;
4.2.1.9 | Universal, hard and/or soft wedging providing wedge angles of up to 60° (in the direction of the MLC movement);
4.2.1.10 | Maximum photon beam field size at the isocentric plane of 40 cm × 40 cm (50% isodose level);
4.2.1.11 | Dose rate up to at least 300 MU/min for photons and electrons;
4.2.1.12 | A low dose rate mode shall be provided for photons;
4.2.1.13 | A light field to indicate the radiation field aperture and a graticule to indicate the principal axes and collimator axis of rotation;
4.2.1.14 | Optical backpointer;
4.2.1.15 | An integrated MLC with at least 80 leaves providing 1 cm leaf width at the isocentre;
4.2.1.16 | A standard non-divergent beam block system with coded (verifiable) accessory trays for mounting; a range of block shapes appropriate to the highest photon energy including eye, spinal cord, triangular and rectangular shapes
4.2.1.17 | A set of at least 4 different electron applicators ranging from approximately 4 cm – 20 cm diameter or side of square. Each applicator shall include a standard set of coded cut-outs (shapes).
4.2.1.18 | A dual independent system of ionization chambers for monitoring dose, dose rate, beam symmetry and beam steering;
4.2.1.19 | An independent backup system to indicate accumulated monitor units and setup parameters at the time of a power failure;
4.2.1.20 | Hand pendants in the treatment room to allow manual control of linac and treatment table movements;
4.2.1.21 | A treatment table with motorized lateral, longitudinal and vertical movements, as well as column rotation (± 180°) and isocentric table rotation up to ± 120°;
4.2.1.22 | The table shall be fitted with a system to enable manual removal of the patient in the event of a power failure;
4.2.1.23 | The table top to be carbon fibre and indexed to allow reproducible placement of immobilization equipment. The table top shall be the same type and model as on the simulator (model tbc) and CT scanner (model tbc);
4.2.1.24 | A computerized control console with slave display inside the treatment room. All the functions and modes of the accelerator must be controlled via software. The console shall have various modes including clinical, physics and service modes. The control console shall have an on/off removable key;
4.2.1.25 | Access to physics and service mode shall only be possible by password and key control;
4.2.1.26 | The linac shall include an integrated amorphous silicon EPID mounted on a motorised arm for planar portal imaging and include software at the treatment control for comparing/matching EPID images with TPS generated DRRs;
4.2.1.27 | The EPID software shall be capable of image enhancement, overlay and matching tools, electronic image approval/rejection, editing and commenting. The software shall be capable of providing the user with suggested couch movements to achieve an improved setup.
4.2.1.28 | The EPID shall be capable of imaging at least a set field size of 25 cm x 25 cm with collision detection.
4.2.1.29 | The EPID shall be capable of sending images directly to the dosimetry server for routine QC purposes (model tbc). See Appendix 1 for on-site configuration.
4.2.1.30 | A voltage stabilizer/power conditioner;
4.2.1.31 | A water chiller for linac cooling (closed system);
4.2.1.32 | Two colour CCTV systems for viewing of the patient from the console (one with pan, tilt and zoom) and a bidirectional audio intercom system;
4.2.1.33 | Emergency stop buttons shall be located on the table, treatment console and around the treatment room.
4.2.1.34 | Wall-mounted laser system indicating the position of the isocentre including 2 side lasers (vertical and longitudinal) and 1 sagittal laser.
4.2.2 | Record and verify system (RVS) consisting of:
4.2.2.1 | A remote server with UPS including an automated daily backup system to an external hard drive or equivalent system; auto-detect and auto-shutdown after 20 minutes in the event of a power failure. The server hardware shall be capable of supporting multiple additional software licences and at least 5 more workstations into the future without a loss in performance and speed.
4.2.2.2 | The RVS shall form the core of the departmental intranet and shall not be connected to the internet (see Appendix 1 for equipment layout). Software installation and upgrades/updates shall be performed using CD/DVD and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection. The RVS shall be provided with advanced anti-virus software which is automatically updated during preventative maintenance by an authorised field service engineer.
4.2.2.3 | The RVS shall be HL7 compliant for future gateway to the HIS for patient administrative fields only (retrieve only). Mandatory fields shall be used to ensure and internally validate unique patient ID, e.g. first name, surname, gender, date of birth and national ID number.
4.2.2.4 | Workstations:
4.2.2.4.1 | Direct interface to the linac control console such that patient treatment data is automatically downloaded, including an in-room system. The system shall be supported by a local UPS such that there is no loss of data (e.g. current MLC leaf position, delivered MU, etc.) in the event of a power failure to the linac. A local shutdown to the workstation shall then be possible.
4.2.2.4.2 | Three workstations in the treatment planning area with:
4.2.2.4.2.1 | Direct interface to the 3D treatment planning system (model tbc) for transfer of treatment parameters and DRRs including a validity check of unique patient administrative information;
4.2.2.4.2.2 | Direct interface to the simulator workstation (model tbc) for transfer of simulation images and treatment parameters including a validity check of unique patient administrative information;
4.2.2.4.2.3 | Direct interface to the virtual simulation workstation (model tbc) for transfer of DRRs and treatment parameters including a validity check of unique patient administrative information;
4.2.2.4.2.4 | A workstation for independent electronic chart check and approval in the medical physics office. This workstation shall be equipped with a local laser printer.
4.2.2.5 | Software features to include:
4.2.2.5.1 | Digital photos of patient (ID and/or setup photos);
4.2.2.5.2 | Automated logging of cumulative dose;
4.2.2.5.3 | Free text entry of setup instructions or alerts based on cumulative dose;
4.2.2.5.4 | Hierarchical security features including requirement for authorised approval of the dose prescription and field parameters prior to treatment;
4.2.2.5.5 | Complete log of activities and users;
4.2.2.5.6 | Patient appointments;
4.2.2.5.7 | Generation of statistical data according to user-defined fields, e.g. diagnosis and managing consultant.;
4.2.2.5.8 | Library of diagnoses according to ICD-10;
4.2.2.5.9 | Ability to correctly log cumulative dose in the event of a treatment interruption or termination.
4.2.3 | The linac shall be capable of upgrade to:
4.2.3.1 | IMRT treatments including step and shoot as well as volumetric arc therapy;
4.2.3.2 | An integrated kilovoltage x-ray imaging system for 2D planar imaging;
4.2.3.3 | 3D kV CBCT imaging of the patient in the treatment position including software for comparing images to planning images and for calculating couch shifts;
4.2.3.4 | EPID Portal Dosimetry.
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
15 | Uptime and Penalties
15.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
15.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
15.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
15.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 4 (hidden)
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 4 - High Dose Rate Brachytherapy Equipment Package
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | International Electrotechnical Commission, Medical Electrical Equipment, Part 1: General requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012, IEC, Geneva (2012)
2.2 | International Electrotechnical Commission, Medical Electrical Equipment, Part 2-11: Particular requirements for the safety of remote controlled automatically-driven Gamma-ray afterloading equipment, Rep. IEC 60601-2-17:2013, IEC, Geneva (2013)
2.3 | International Electrotechnical Commission, Medical Electrical Equipment: Requirements for the safety of radiotherapy treatment planning systems, Rep. IEC 62083:2009, IEC, Geneva (2009)
2.4 | International Organization for Standardization, Sealed Radioactive Sources- General requirements and classification, ISO 2919:2012, ISO, Geneva (2012)
2.5 | International Organization for Standardization, Sealed Radioactive Sources- Leakage test methods, ISO/TR 4826, ISO, Geneva (1979)
2.6 | “Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards”, IAEA, Vienna (2014)
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | a high dose rate brachytherapy remote afterloading treatment unit with a dedicated treatment planning system is required for treatment of cancer patients with brachytherapy;
4.1.2 | a patient treatment couch is required for support of the patient at the time of treatment and simulation;
4.1.3 | applicators and accessories are required for proper placement of the brachytherapy source at various anatomical sites;
4.1.4 | ancillary equipment for safety and quality control;
4.1.5 | table top processor; and
4.1.6 | the System shall interface with the C-arm (Siemens Cios Connect) in the HDR treatment bunker, simulator (model tbc) and CT scanner (model tbc). Equipment configuration is shown in Appendix 1.
4.2 | Technical Requirements
4.2.1 | At least 10 channels remote controlled afterloading system for intracavitary high dose rate brachytherapy capable of operating with a Co-60 or Ir-192 microsource;
4.2.2 | Precise double drives for a miniaturized source and a dummy source with continuous digital positioning including pre-exposure routine dummy test run for safety and catheter length check;
4.2.3 | A control unit, which facilitates operation, controlling and monitoring functions including:
4.2.3.1 | A connection to the dedicated treatment planning system via a network;
4.2.3.2 | Password-defined access for all users with restricted access to:
4.2.3.2.1 | data of high security, e.g. source strength;
4.2.3.2.2 | a log of all users’ activities including patient identification, plan identification, channel numbers, source positions, source dwell times and the number of source transfers;
4.2.3.2.3 | a database (library) of uniquely defined standard treatment plans for rigid applicator sets and line sources with automatic update of source dwell times. Access to the database for plan editing purposes shall be password-controlled and restricted;
4.2.3.3 | Unique patient identification and treatment plan preparation;
4.2.3.4 | Display of status information including during a power interruption;
4.2.3.5 | Online recording of treatment with visual graphic display of treatment progress;
4.2.3.6 | Hardcopy output of the entire treatment protocol after treatment (A4 laser printer to be supplied);
4.2.4 | The treatment afterloader unit shall automatically verify all unique applicator connections and be easily manoeuvrable within the treatment room;
4.2.5 | Internal battery or UPS to ensure safe source or dummy retraction in the event of an emergency including a crank for manual source retraction, long handle forceps and an emergency container thick enough for storage of a new Co-60 or Ir-192 HDR source and large enough to fit the largest possible applicator;
4.2.6 | A patient treatment table including:
4.2.6.1 | Head plate, removable leg plates, stirrups, seat plate extension and backrest manoeuvrability in order to treat a patient in the seated position;
4.2.6.2 | The tabletop to be radiolucent to enable X-ray imaging;
4.2.6.3 | The table design to be such that a C-arm (model Cios Connect) can easily be positioned and moved around the patient to take images of the pelvis from any angle.
4.2.7 | Applicators and accessories including:
4.2.7.1 | 6 complete sets of CT compatible IU-Ring Applicator sets (IU angle 30° from applicator axis) with rectal retractors; at least 3 different ring diameters and 3 different IU lengths;
4.2.7.2 | 3 complete sets of CT compatible, straight cylindrical vaginal applicators including at least 3 different diameters;
4.2.7.3 | 4 sets of accessories for each of the above applicator sets (i.e. transfer tubes, X ray markers, ring caps);
4.2.7.4 | 6 Sterilization Boxes appropriate to the above range of applicator sets;
4.2.7.5 | 1 of Reconstruction box for use in gynaecological applications.
4.2.8 | Brachytherapy treatment planning system capable of accepting DICOM images from the C-arm, simulator and CT scanner (models tbc) including:
4.2.8.1 | Ability to plan for either a Co-60 or Ir-192 HDR microsource;
4.2.8.2 | Hierarchical password-protected access to sensitive data, e.g. source calibration, approval of volumes and treatment plans, etc.;
4.2.8.3 | Automatic internal check and update of source strength in conjunction with the treatment console;
4.2.8.4 | Library of applicators;
4.2.8.5 | Automated reconstruction of applicators from non-isocentric C-arm or isocentric simulator images or CT study using key reference points;
4.2.8.6 | Optimisation of source dwell positions and/or dwell times based on prescribed dose to a target volume/points and constraints to organ at risk volumes/points;
4.2.8.7 | The formalism of the AAPM TG43 (and all supplements as relevant) shall be used for the calculation of absorbed dos;
4.2.9 | The console and treatment planning system shall be equipped with an automatic overnight backup system to an external hard drive.
4.2.10 | The HDR and treatment planning system shall form part of the End-User’s intranet and shall not have access to the internet. Software installation and upgrades/updates shall be performed using CD/DVD and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection.
4.2.11 | The HDR afterloader and treatment console shall be fitted with a rechargeable power supply so that a typical treatment can be completed in the event of a power failure including the functionality of all safety devices.
4.2.12 | Ancillary equipment including:
4.2.12.1 | Portable survey meter to measure potential contamination on the source during leak test and to verify source retraction;
4.2.12.2 | Source position check ruler;
4.2.12.3 | Table-top processor (for wet film);
4.2.12.4 | Radiation area monitoring system with audible and visible signal for treatment room and control panel monitoring;
4.2.12.5 | radiation warning lights and signage;
4.2.12.6 | door interlocks;
4.2.12.7 | patient intercommunication device;
4.2.12.8 | 2 complete CCTV systems (one with pan, zoom and tilt facility).
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
15 | Sources and Source Management:
15.1 | At the request of the End-User, the Contractor shall provide the following sources and their subsequent management:
15.2 | Six-year source and management contract for Ir-192 micro-sources including 4 source changes per year;
15.3 | Six-year source and management contract for Co-60 micro-sources including a source change every 3 years;
16 | Uptime and Penalties
16.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
16.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
16.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
16.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 5 (hidden)
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 5 - Radiotherapy Treatment Planning system
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Radiotherapy Equipment - Co-ordinates, Movements and Scales, Rep. IEC 61217, IEC, Geneva (1996).
2.2 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, General requirements for safety 4. Collateral standard: Programmable electrical systems, Rep. IEC 60601-1-4 Part 1, IEC, Geneva (1996).
2.3 | INTERNATIONAL ELECTROTECHNICAL COMMISSION, Medical electrical equipment – Requirements for the safety of radiotherapy treatment planning systems, Rep. IEC 62083, IEC, Geneva (2000).
2.4 | IAEA, Commissioning of Radiotherapy Treatment Planning Systems: Testing for Typical External Beam Treatment Techniques, IAEATECDOC-1583, Vienna (2008).
2.5 | IAEA, Specification and Acceptance Testing of Radiotherapy Treatment Planning Systems, IAEA-TECDOC-1540, Vienna (2007).
2.6 | “Radiation Protection and Safety of Radiation Sources: International Basic Safety Standards”, IAEA, Vienna (2014)
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | The TPS shall meet the requirement for 3D CRT planning of radiotherapy treatments from patient CT image data sets including those from virtual simulation.
4.1.2 | The System shall be capable of transferring DRRs and treatment parameters directly to the RVS. Beam data input and commissioning shall be facilitated through direct input from the dosimetry server.
4.2 | Technical Requirements
4.2.1 | A server and three (two treatment planning and one contouring) workstations including:
4.2.1.1 | dual 19-inch monitors;
4.2.1.2 | the server shall be equipped with an automatic overnight backup system to an external hard drive.
4.2.1.3 | Direct connectivity to the End-User’s PACs system is required for long-term image archive.
4.2.1.4 | The System shall form part of the End-User’s RVS intranet and shall not have access to the internet.
4.2.1.5 | Software installation and upgrades/updates shall be performed using CD/DVD and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection.
4.2.1.6 | On-line UPSs to for the server and each workstation for at least thirty minutes in order to facilitate controlled shutdown in the event of a power failure
4.2.2 | Ability to import patient CT data sets from the scanner using the DICOM3 standard (model tbc); the System shall distinguish between contrast (scans to enhance target visualisation not necessarily in the treatment position) and non-contrast scans (planning scans in the treatment position) of the same patient;
4.2.3 | Ability to interface with a CT simulation workstation (model tbc) for the calculation of treatment times through the importation of CT studies with volumes, isocentre localisation, field placement and field design (shielding blocks or MLC) of external beam radiotherapy fields (see Appendix 1 for installation configuration), including generation of digitally reconstructed radiographs;
4.2.4 | Ability to interface with a simulation workstation (model tbc) for sending DRRs or receiving simulation films to upload onto a water-equivalent ICRU phantom with user-defined dimensions.
4.2.5 | 3D visualisation capability for patient data display, beam display, room’s eye view and dose display; multi-planar reconstruction views (major axes and user-defined)
4.2.6 | Contouring tools which allow the definition in 3D of structures including target, organs at risk and patient outline;
4.2.7 | Automated tools to allow:
4.2.7.1 | registration of additional image sets for a single patient
4.2.7.2 | the expansion of the CTV to a PTV and organ at risk volumes to PRVs with non-uniform margins in three dimensions;
4.2.8 | Ability to add bolus structures to the patient data set of various shape and density;
4.2.9 | A comprehensive “forward planning” environment which allows the planner to modify beam weights, beam positioning, wedges and blocks to optimise the treatment plan; the ability to generate a library of dose plans shall be available;
4.2.10 | Modern, well-validated photon and electron algorithms which calculate the dose to the patient considering the 3D nature and heterogeneity of the patient data set;
4.2.11 | Ability to override density values (artefacts);
4.2.12 | Ability to allow dose prescription to a point, volume or isodose level with user-defined normalisation;
4.2.13 | 4.2.6 statistics, and 2D and 3D dose visualisation including plan comparison tools;
4.2.14 | Ability to generate, view and export DRRs;
4.2.15 | A comprehensive beam modelling module which allows the creation of complete geometric and dosimetric models for the Linac beams. The module shall have the following features:
4.2.15.1 | Ability to import measured beam profiles and output factors,
4.2.15.2 | Ability to model dynamic, fixed and internal wedges,
4.2.15.3 | Tools to allow the comparison of the beam model and measured data from the dosimetry server (model tbc),
4.2.15.4 | Security features which protect beam data and beam models from modification;
4.2.15.5 | User and password security which allows approval/locking of volumes, treatment plans and different levels of access to the functionality of the TPS based on the user’s profile e.g. administrator, physicist, senior planner, planner;
4.2.16 | A networked colour A3 laser printer for printing of treatment plan information and isodose plots with patient images; toner for one year shall be provided.
4.2.17 | The server hardware shall be capable of supporting multiple additional software licences and at least 2 more workstations into the future without a loss in performance and speed including:
4.2.17.1 | inverse planning for IMRT (static and arc)
4.2.17.2 | more advanced planning algorithms,
4.2.17.3 | a software module for dosimetric verification of treatment plans based on EPID dosimetry.
4.2.18 | The TPS shall have capability of integration with the RVS (model tbc) at the Site for exporting patient administration data, beam delivery data and DRRs.
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
15 | Uptime and Penalties
15.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
15.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
15.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
15.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 6
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number: | 160277-2
Lot 6 - Radiotherapy Dosimetry Equipment
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Applicable Documents
2.1 | IEC 60731:2011 Medical electrical equipment - Dosimeters with ionization chambers as used in radiotherapy.
4 | Requirements
4.1 | Functional and Perfomance Requirements
4.1.1 | A complete 3D beam data acquisition system for clinical commissioning of external beam radiotherapy equipment as well as performing routine quality control measurements;
4.1.2 | Various items of dosimetry equipment to perform reference dosimetry, relative dosimetry and in-vivo dosimetry.
4.1.3 | Accessories to support quality dosimetry measurements.
4.1.4 | A dosimetry server to support analysis, patient-specific QA measurements and to act as a central database for dosimetry measurements
4.1.5 | Devices for performing mechanical, protection and safety checks on radiotherapy equipment.
4.1.6 | Dosimetry and quality control equipment for the radiotherapy imaging equipment
4.2 | Technical Requirements
4.2.1 | The 3D beam data acquisition system shall meet the following technical requirements:
4.2.1.1 | Acrylic water tank with dimensions such that at least a water volume of 50 x 50 x 50 cm3 can be scanned;
4.2.1.2 | Detector position accuracy: ± 0.1 mm and position reproducibility: ± 0.1 mm;
4.2.1.3 | Water tank with programmable motorized scanning capability in the X, Y and Z planes for ionization chambers or diode detectors;
4.2.1.4 | A lift table with vertical travel range of at least 500 mm and a tilt function to allow for accurate levelling of the scanning arms;
4.2.1.5 | A dual electrometer system with bias range 0 through 400 V, minimum resolution of 10 fA, and leakage current < 250 fA. The system shall be capable of measurements with or without a reference detector;
4.2.1.6 | The control unit is to control movement of moving mechanisms and interface with electrometer;
4.2.1.7 | Appropriate cables to provide connectivity between the reservoir, watertank, control unit and electrometer system;
4.2.1.8 | A water reservoir with bidirectional water pump to the water tank and volume capacity of at least 150 litres;
4.2.1.9 | One portable laptop computer with connectivity to the control unit;
4.2.1.10 | Software for data acquisition with scan optimization, data handling and analysis, and a software license for automated transfer of data to the Treatment Planning System (model tbc) format;
4.2.1.11 | The software shall allow offline export, smoothing, analysis, PDD/TMR conversion, ionization to dose conversion, isodose generation and manipulation of beam profile data or depth dose data in Excel format. Additional licenses may be necessary to enable data analysis on the dosimetry server;
4.2.1.12 | Two thimble waterproof ionization chambers with approximate sensitive volume of approximately 0.13 cm3 with TNC connectors. Metallic build-up caps for 6 MV and 10/15 MV in air measurements;
4.2.1.13 | One un-shielded waterproof diode with TNC connectors;
4.2.1.14 | One waterproof parallel plate electron chamber;
4.2.1.15 | Holders for above detectors;
4.2.1.16 | A holder for a waterproof Farmer chamber;
4.2.1.17 | 1D array for measurements of soft wedges;
4.2.1.18 | Connection cables between the computer and controller (permanently installed);
4.2.1.19 | Two TNC cables for connectivity between the scanning detectors and the dual electrometer system (may be permanently installed).
4.2.2 | The dosimetry equipment to perform reference dosimetry, relative dosimetry and in-vivo dosimetry shall meet the following technical requirements. For the applicable items, the Contractor shall provide DKD calibration, or equivalent. If not DKD, the Contractor shall give details of the calibration.
4.2.2.1 | Universal electrometer with DKD calibration with TNC connector capable of reverse polarity and 0 – 400 V measurements;
4.2.2.2 | Two waterproof Farmer type chamber 0.6 cm³, DKD calibrated in terms of absorbed dose to water at Co-60 with TNC connector;
4.2.2.3 | A parallel plate electron ionisation chamber, DKD calibrated in terms of absorbed dose to water at Co-60 with TNC connector;
4.2.2.4 | An ionisation chamber for use in general radiography/fluoroscopy, DKD calibrated in terms of RQR beam qualities with TNC connector;
4.2.2.5 | 1D water tank with chamber holder for Farmer type and electron ionisation chambers;
4.2.2.6 | Two 25 m low noise extension cable with TNC connectors;
4.2.2.7 | Well chamber with electrometer for high dose rate brachytherapy with DKD calibration for Co-60 and Ir-192 sources;
4.2.2.8 | CTDI 16/32 cm diameter phantom and pencil ionisation chamber, DKD calibrated in terms of RQT beam qualities (TNC connectors are preferred);
4.2.2.9 | In-vivo diode dosimetry system (complete with at least a 5-channel electrometer and permanent cabling) for organ at risk/superficial/skin, Co-60, 6 MV, 10/15 MV and electron diodes. At least 2 diodes per energy are required. Dosimetry results shall be directly transferrable to the electronic file of the patient in the record and verify system (model tbc). A calibration phantom would be highly recommended.
4.2.3 | Required Accessories to support quality dosimetry measurements shall meet the following technical requirements:
4.2.3.1 | precision barometer for measurement of absolute air pressure with DKD calibration;
4.2.3.2 | precision thermometer for measurement of ambient temperature with DKD calibration;
4.2.3.3 | Catphan 600 phantom for CT quality control and treatment planning system commissioning measurements;
4.2.3.4 | EPID quality control phantom including software for automated analysis.
4.2.4 | Dosimetry server shall meet the following technical requirements:
4.2.4.1 | Desktop computer including operating system, Microsoft Office and all dosimetry analysis software to enable offline analysis and data manipulation
4.2.4.2 | Flat bed scanner and A4 laser printer;
4.2.4.3 | Planar image/film analysis and calibration software;
4.2.4.4 | Direct connectivity to the beam acquisition laptop, treatment planning system to allow data transfer (model tbc), CT scanner (model tbc), to the external beam treatment unit to allow EPID image transfer (model tbc) and scanner to allow automated 1D or 2D analysis of film;
4.2.4.5 | Hardware and performance shall adequate to accept future installation of additional dosimetry software packages, e.g. IMRT verification software;
4.2.4.6 | External hard drive backup system and 30 minutes UPS to allow controlled shut down in the event of a power failure;
4.2.4.7 | The server shall form part of the End-User’s intranet and shall not have access to the internet. Software installation and upgrades/updates shall be performed using CD/DVD and all other external data device inputs shall be disabled to enhance data protection and minimise virus infection.
4.2.5 | Devices for performing mechanical, protection and safety checks shall meet the following technical requirements:
4.2.5.1 | Radiation survey meter with DKD calibration appropriate to imaging and high energy therapy equipment;
4.2.5.2 | Neutron survey meter with DKD calibration appropriate to high energy photon beams;
4.2.5.3 | Two calibrated personal audible dosimeters;
4.2.5.4 | Spirit level for levelling of water phantom, confirming gantry scales, etc.
5 | Site readiness (Bidder shall address this in the implementation plan)
5.1 | After contract award, the Contractor shall visit the Site to determine readiness. If modification of the End User’s Site for the System is required, the Contractor shall verify the final technical drawings prepared by the End-User and certify that they conform to the Contractor’s technical requirements prior to commencement of the modification and/or construction of the premises for the System/Equipment. The Contractor shall arrange for the visit which includes travel, accommodation, and allowance. The Contractor shall bear all costs associated with such visit.
5.2 | The Contractor shall communicate with the End-User’s coordinator and verify the appropriateness of the designated area for the installation of the System. Furthermore, the Contractor shall inform on time the End-User of any additional requirement (structural, electrical, IT infrastructure) necessary for the installation of the System.
5.3 | The completion date of the Site preparation will be communicated by the IAEA to the Contractor in due time to start the execution of the Contract activities.
5.4 | Notwithstanding any authorisation given by the IAEA, the Contractor shall visit, inspect and ascertain that all necessary conditions are met at the End-User’s Site before starting any activities. Any comments or suggestion as regard the conditions of the Site shall be made at least four (4) weeks before initiating the installation activities.
5.5 | Appendix 1 provides the site layout and indicates the preferred positions of the equipment and workstations.
6 | Licenses and Permits
6.1 | In addition to the provisions of Article 3.1 (“Responsibilities of the Contractor”) and Article 27.2 (“Export Licenses”) of the General Conditions of Contract, the Contractor shall obtain and maintain all permits and licenses necessary for the supply and delivery of the System to the End-User. The Contractor shall timely provide the End User with all information and technical support that may be needed to obtain the import licenses, permits and/or authorizations in accordance with the regulations of the appropriate regulatory authorities.
7 | Marking
7.1 | The System shall have all safety markings in the English language.
8 | Packing
8.1 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
9 | Quality Requirements
9.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
9.2 | The Contractor shall document the compliance with this quality assurance system.
10 | Testing and Acceptance
10.1 | The System, prior to shipment, shall be tested for conformance of the System with manufacturer’s performance specifications and the minimum requirements specified herein.
10.2 | The System, after installation, shall be tested by the Contractor together with the End-User to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified herein as determined by the IAEA and the End-User.
10.3 | The results of the testing of the System shall be documented by the Contractor in an acceptance protocol that shall be signed by the End-User.
11 | Installation and Training
11.1 | The Contractor shall install the System at the Site.
11.2 | The Contractor shall provide the End-User a list of the equipment and tools to be used during installation and to be re-exported at the end of the work, eight (8) weeks in advance of the expected importation date to allow the End-User to obtain the necessary import and export permits.
11.3 | No installation activities shall be initiated unless authorised by the IAEA.
11.4 | Provided that the Site is ready and IAEA has given proper authorisation, the Contractor, in agreement with the End-User, shall install the System at the Site. The Contractor shall provide for this purpose all necessary tools and staff (including their travel, accommodation and subsistence as necessary).
11.5 | The Contractor shall provide at least five (5) days training for up to six (6) staff of the End-User in the operation and maintenance of the System at the Site immediately after the installation of the System.
12 | Deliverable Data Items
12.1 | The Contractor shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English language.
13 | Warranty
13.1 | The System shall be covered by one (1) year warranty, including parts and labour, starting as of the date of successful on-site acceptance, as per Section 10 above. Warranty shall also cover hardware and software upgrades and updates.
13.2 | Warranty shall include all necessary spare parts, shipment to site, cost of replacement (work, personnel etc.) and disposal of faulty parts.
14 | Maintenance and Spare Parts
14.1 | The Contractor shall provide on-site full maintenance services during warranty period, for the proper functioning of the System. Full maintenance services during the warranty period shall include:
14.1.1 | preventative maintenance
14.1.2 | on-call interventions
14.1.3 | any safety, software and hardware update and upgrade for the System that will become available
14.1.4 | all necessary replacement and spare parts
14.2 | As part of the on-site acceptance, the Contractor shall provide to the local hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention.
14.3 | The Contractor shall provide evidence of the capability to adequately provide technical support for the System in the future, in a timely manner, stating the network of official representatives in the country and/or in the region.
14.4 | Intervention time shall be clearly defined and shall comply with the uptime requirements defined in section 15 below.
14.5 | Upon installation and without prejudice for warranty obligations of the Contractor, an initial set of essential spare parts shall be provided to be stored at the Site. A list of available spare parts and prices shall subsequently be provided and updated as necessary.
14.6 | At the request of the End-User, the Contractor shall provide five (5) additional years of full maintenance services (as defined in section 14.1 above), following the initial one (1) year full warranty.
14.7 | At the request of the End-User, the Contractor shall perform a DKD re-calibration every two (2) years for the reference chambers 0.6cc electron diagnostic and well chamber.
15 | Uptime and Penalties
15.1 | The Contractor guarantees that the System shall have an up-time of at least 95% (excluding outages for scheduled maintenance or causes external to the System).
15.2 | Uptime is calculated on a basis of two hundred and fifty (250) operating days per year (weekly working days).
15.3 | Should the down time exceed two (2) working days cumulative on a six-months basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period.
15.4 | The records of downtime of the System will be kept by a representative of the End-User, at the Site. The Contractor shall have the right to request copies of such records.

## Lot 7 (hidden)
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specification.
RFQ Number:
Lot 7 - Immobilisation Devices
Specification section | Description | Compliance Y/N | Evidence (state the evidence and indicate the technical proposal page number) | Bidder's Comment
1 | Scope | -
2 | Requirements
2.1 | The Contractor shall provide the Equipment indicated in the attached list (Table 1), or alternatives.
3 | Marking and Packing
3.1 | The System shall have all safety markings in the English language.
3.2 | The System, for the shipment by air/road to the End-User, shall be packed in accordance with international standards that are applicable for the shipment by air/road of this kind of equipment.
4 | Quality Requirements
4.1 | The System shall be manufactured, shipped and installed in accordance with the Contractor’s ISO quality assurance system or an equivalent quality assurance system.
4.2 | The Contractor shall document the compliance with this quality assurance system.
Table 1
Item | Product Commodity | Item description | Quantity | Unit of Measure | Compliance Y/N
1 | Immobilization device | Two-pin indexing bars for CT, simulator, linac tabletop and mould room. Compatible with carbon fibre couch tops (model tbc). | 4 | each
2 | Immobilization device | Superflab/jelly Bolus, 30cm X 30 cm X 0.5 cm | 5 | sheets
3 | Immobilization device | Superflab/jelly Bolus, 30 cm X 30 cm X 2.5 cm | 5 | sheets
4 | Immobilization device | Superflab/jelly Bolus, 30 cm X 30 cm X 1.5 cm | 5 | sheets
5 | Immobilization device | Thermoplastic Bolus, 30 cm X 30 cm X 0.5 cm | 15 | sheets
6 | Medical Equipment | Large Waterbath with Digital temperature control, 220V/50Hz | 1 | each
7 | Immobilization device | Adult Head and neck baseplate suitable for 3 and 5 point masks. | 3 | each
8 | Immobilization device | Paediatric Head and neck baseplate suitable for 3 and 5 point masks. | 3 | each
9 | Immobilization device | Thermoplastic perforated head masks, 3 point (adult) | 5 | box of 10
10 | Immobilization device | Thermoplastic perforated head masks, 3 point (paediatric); pre-cut eye and mouth openings | 1 | box of 10
11 | Immobilization device | Thermoplastic perforated head and shoulder masks, 5 point (adult) | 5 | box of 10
12 | Immobilization device | Thermoplastic perforated head and shoulder masks, 5 point (paediatric); pre-cut eye and mouth openings | 1 | box of 10
13 | Immobilization device | Supine headrests (adult), rigid or high density foam; neutral, extended and flexed positions | 8 | sets
14 | Immobilization device | Supine headrests (paediatric), rigid or high density foam; neutral, extended and flexed positions | 4 | sets
15 | Immobilization device | Carbon fibre rectangular block and wedge extenders compatible with headrests (adult) | 4 | sets
16 | Immobilization device | Carbon fibre rectangular block and wedge extenders compatible with headrests (paediatric) | 4 | sets
17 | Immobilization device | Carbon fibre breast board with indexed bi-lateral and single arm support including hand grips and head rest. Maximum board tilt of at least 250; buttock or knee rest. | 4 | each
18 | Immobilization device | Indexable baseplate compatible with carbon fibre couch tops (model tbc) including adjustable, indexable high density foam knee and foot rests | 4 | each